# ASTRA 6 Report Experiment7

File Name: Experiment7 
Collection Operator: AYRESGPC\Ayres Research Group (AYRESGPC\Ayres Research Group (Ayres Research Group)) 
Processing Operator: AYRESGPC\Ayres Research Group (Ayres Research Group)

### Sample: sample

## Configuration

Concentration Source: RI 
Flow Rate: 1.000 mL/min

### Light Scattering Instrument: miniDAWN TREOS

Cell Type: Fused Silica 
Wavelength: 656.0 nm 
Calibration Constant: 6.1050×10-5 1/(V cm)

### RI Instrument: Optilab rEX

### Viscometer: ViscoStar

Dilution Factor: 0.4976

### Solvent: thf

Refractive Index: 1.402

## Processing

Collection Time: Friday January 31, 2014 03:06:45 PM Eastern Daylight Time 
Processing time: Friday January 31, 2014 03:41:46.437 PM Eastern Daylight Time

### Peak settings:

| Peak Name | Peak 1 | Peak 2 | Peak 3 |
| --- | --- | --- | --- |
| Light Scattering Model | Zimm | Zimm | Zimm |
| Fit Degree | 1 | 1 | 1 |
| dn/dc (mL/g) | 0.1850 | 0.1850 | 0.1850 |
| A2 (mol mL/g²) | 0.000 | 0.000 | 0.000 |
| Injected Mass (µg) | 100.000 | 255.000 | 475.000 |
| Viscometry Model | Huggins | Huggins | Huggins |
| Huggins Equation Parameter | 0 | 0 | 0 |
| Kraemers Equation Parameter | 0 | 0 | 0 |

### Results Fitting Procedure:

| Data | Fit Model | Degree | R2 | Extrapolation |
| --- | --- | --- | --- | --- |

## Results

| Peak Results |  |  |  |
| --- | --- | --- | --- |
|  | Peak 1 | Peak 2 | Peak 3 |
| Masses |  |  |  |
| Calculated Mass (µg) | 59.69 | 151.02 | 237.80 |
| Molar mass moments (g/mol) |  |  |  |
| Mn | 6.157×105 (±0.206%) | 1.933×105 (±0.638%) | 2.812×104 (±3.533%) |
| Mp | 6.461×105 (±0.151%) | 2.082×105 (±0.511%) | 2.966×104 (±2.797%) |
| Mv | 6.718×105 (±0.071%) | 1.951×105 (±0.082%) | 2.840×104 (±0.502%) |
| Mw | 6.865×105 (±0.228%) | 1.953×105 (±0.562%) | 2.845×104 (±3.192%) |
| Mz | 8.417×105 (±0.649%) | 1.970×105 (±1.209%) | 2.874×104 (±6.863%) |
| Polydispersity |  |  |  |
| Mw/Mn | 1.115 (±0.308%) | 1.010 (±0.850%) | 1.012 (±4.761%) |
| Mz/Mn | 1.367 (±0.681%) | 1.019 (±1.367%) | 1.022 (±7.719%) |
| rms radius moments (nm) |  |  |  |
| Rn | 30.4 (±0.7%) | 2.5 (±301.7%) | n/a |
| Rw | 33.2 (±0.6%) | 1.8 (±542.5%) | n/a |
| Rz | 38.9 (±0.5%) | 1.0 (±1793.5%) | n/a |
| Intrinsic viscosity moments (mL/g) |  |  |  |
| ηn | 168.677 (±2.158%) | 72.048 (±1.515%) | 18.603 (±3.076%) |
| ηw | 180.89 (±2.84%) | 71.96 (±1.36%) | 18.49 (±2.97%) |
| ηz | 208.373 (±4.435%) | 71.870 (±1.241%) | 18.386 (±2.869%) |
| Hydrodynamic radius moments (nm) |  |  |  |
| Rh(n) | 25.3 (±0.7%) | 13.0 (±0.5%) | 4.3 (±1.6%) |
| Rh(w) | 26.7 (±0.9%) | 13.0 (±0.5%) | 4.4 (±1.5%) |
| Rh(z) | 29.6 (±1.3%) | 13.1 (±0.5%) | 4.4 (±1.5%) |

Mark-Houwink-Sakurada a: 0.719 (±0.145%) 
Mark-Houwink-Sakurada K: 1.132×10-2 (±1.297%) mL/g