## Slide 1: Expanded Light Scattering Flow Cell Assembly

## Slide 2: Light Scattering Flow Cell Assembly Parts

## Slide 3
Manifold Disassembly
A
B
C
D
1
3
2

## Slide 4
Window Removal
Window retainer
removal tool
Flow cell window o-ring

## Slide 5: 1st Critical Step: Flow Cell Window Cleaning
Look through window towards edge of bright light source and move between and area of high and low contrast
Use loupe (4X or greater) to inspect critical surfaces.
Look for chips, scratches, fibers and spots.
Fibers/Scratches
Water/Salt Spots
Chips
little
big

## Slide 6: Pre-Stretching of O-rings
~30 minutes to relax back to original size
Window retainer
removal tool
Plump
New
Old
Flat
0.02 µm filtered
alcohol
1
2
3
vs.

## Slide 7: Flow Cell Window Re-Installation
Righty tighty; Lefty loosy
Hand tight is just right.
Easiest to grip FC window side-on for re-install to the manifold
Flow cell window o-ring

## Slide 8: 2nd Critical Step: Flow Cell Glass Cleaning
Critical Optical Surfaces:
A - Inner Bore
B - Rounded Surfaces
Finger print/chip
BAD
GOOD
Finger print
Small Chip
Smooth
Reflection
B
A

## Slide 9: 3rd Critical Step: Re-assemble flow cell and confirm alignment
GOOD - centered bore
BAD
GOOD - concentric
BAD
Alignment Clues:
“Step on Stops”

## Slide 10
Summary: Three “Critical” Steps
1st Critical Step: Confirm clean flow cell windows.
2nd Critical Step: Confirm clean flow cell glass.
3rd Critical Step: Confirm alignment after re-assembly.