## Slide 1: Notes on Batch Light Scattering Experiments
Light Scattering University
Wyatt Technology Corporation
Santa Barbara, California
Notes: Notes:
Additional information for each slide is presented in this section.
Please feel free to use this space to take notes.

## Slide 2: Measurements on Unfractionated Samples
Obtain:
Weight-averaged molar mass:
z-averaged mean square radius:
z-averaged rms radius:
Notes: Notes:
For fractionated or monodisperse samples, the measured molar mass and rms radius values from light scattering measurements need no further interpretation.
In measurements on unfractionated samples, however, averaged quantities are obtained: the weight-averaged molar mass, MW, and the z-averaged rms radius, rz.
These averages and different distributions for fractionated vs. unfractionated samples will be discussed in more detail in tomorrow’s lecture.
Key Ideas
weight-averaged molar mass MW - Since the measured light scattering is proportional to the molar mass and the weight concentration, a measurement on an unfractionated sample results in a weight-averaged quantity. See the definition in the above slide.
z-averaged rms radius rz - Again, measured light scattering is proportional to the molar mass and the weight concentration, resulting in the average above. The z-averaged radius is technically a higher moment of the radius distribution. For the case of an ideal polymer coil, the molar mass is proportional to the mean square radius and therefore the expression for rz becomes similar to the expression for the z-average molar mass (hence the name z-average radius). The “z” comes originally from the German word Zentrifuge, since the distribution is similar to that obtained from a centrifuge.

## Slide 3: Range of Measurement
Molar Mass M: 500 Da < M <
1x106 Da miniDAWN
1x108 Da DAWN
Radius rg: 10 nm < rg <
50 nm miniDAWN
500 nm DAWN
( = 660 nm)
Notes: Notes:
Both the miniDAWN TREOS, with three angles of measurement, and the DAWN HELEOS, with eighteen angles of measurement, are capable of measuring molecules with molecular weights down to several hundred Daltons and root mean square radii down to 10nm. The 10nm lower limit is for a =660nm laser.
The miniDAWN is capable of measuring root mean square radii up to 50nm, and the DAWN, with the advantage of its many angles of measurement, is capable of measuring root mean square radii up to 500nm.
For both instruments, the upper limits on the radii set the upper limit on the mass. The lower limits on radii and mass are independent of each other.
Not covered in today’s lecture is Quasi-Elastic Light Scattering (QELS). Wyatt QELS may be integrated with either the miniDAWN or the DAWN, and is capable of measuring the hydrodynamic radius Rh between 2nm and 30nm in flow mode, and 2nm and 1000nm in static mode.

## Slide 4: Light Scattering Data Analysis – Physical Picture
Why fit to:
Zero concentration:
Eliminate intermolecular interactions.
Ideal experiment: Measure scattering from one molecule.
Zero angle:
Eliminate intramolecular interference.
Ideal experiment: Measure scattering at zero angle.
Rewrite Light Scattering Equation using Zimm formalism:
Notes: Notes:
To determine the molar mass of a polymer one would like to study the scattering from:
  a single polymer molecule (infinite dilution) to avoid intermolecular effects
  a detector at zero scattering angle to avoid intramolecular effects.
Since we can achieve neither of the above experimentally, we measure the scattering at multiple angles and several different concentrations.
We can then extrapolate the scattering data to zero angle and zero concentration to determine the absolute molar mass of the polymer.
ASTRA V uses a global fit to fit all data to the Light Scattering equation.

## Slide 5: Plots
Zimm Plot (batch): Measure light scattering at multiple angles and concentrations. Retrieve M, rg, and A2.
Debye Plot (online):
Measure angular variation of light scattering for a single low concentration.
Retrieve M and rg.
Notes: Notes:
For light scattering measurements at different concentrations and angles, it is possible to make a Zimm Plot to determine the molar mass, rms radius, and second virial coefficient.
For low concentrations, such as those obtained after fractionation, it is possible to plot just the angular variation in a Debye Plot to obtain the molar mass and rms radius.
The use of Zimm and Debye plots will be discussed in more detail in tomorrow’s lecture.

## Slide 6: A Zimm Plot - Real Data
200 kDa polystyrene in toluene
Weight average molar mass | (2.019 ± 0.001) e+5 g/mol
Z-average rms radius | 16.7 ± 0.1 nm
Second virial coefficient A2 | (4.100 ± 0.010) e-4 mol mL/g²
 decreasing angle
 decreasing concentration
Notes: Notes:
A good looking Zimm Plot requires accurate concentrations. Concentration differences as small as 1% between the actual concentration and the concentration entered in the software can lead to a very distorted looking Zimm plot.
Q: How many scattering angles and concentrations are required to make a Zimm plot?
A: Enough to accurately fit the angular scattering data to zero scattering angle and concentration.
Although it is true that 2 points determine a straight line, more data points are better. For radii below ~50 nm ( M ~ 1,000,000 for polystyrene in toluene) plotting the reciprocal of the scattering intensity vs. sin squared of the half angle (the Zimm plotting formalism) is linear enough to fit the experimental angular data with a 1st order polynomial. Although the data could be plotted and fit using only 2 angles, the use of 3 angles (as with the mD) results in more confidence and better precision. The use of up to 18 different scattering angles (as with the DAWN) results in even more confidence.
For samples with a polydisperse molar mass distribution and for polymers with radii greater than 50 nm, using even the Zimm plotting formalism may result in curvature. In this case a 2nd order (or higher) polynomial may be required to adequately fit the experimental angular data and allow a reliable fit to zero angle. To plot a 2nd order polynomial more than 3 data points should be used. For polymers with radii greater than 50 nm, therefore, the 18 angle DAWN is recommended.
ASTRA has a “random coil” model option which may be used to analyze polymers with radii greater than 50 nm if the polymer is a random coil. This option extends the range of the miniDAWN.

## Slide 7: A Zimm Plot (Another View)
200 kDa Polystyrene in toluene with negative stretch factor k
Results are unchanged, but plot is easier to view
Weight average molar mass | (2.019 ± 0.001) e+5 g/mol
Z-average rms radius | 16.7 ± 0.1 nm
Second virial coefficient A2 | (4.100 ± 0.010) e-4 mol mL/g²
 decreasing angle
decreasing concentration 
Notes: Notes:
The ASTRA software automatically computes the value of the stretch factor k to properly scale both the angular and the concentration data on the same axis. The user can only change the sign of k.
Changing the sign of the stretch factor in the Zimm plot does not change the computed results. It only changes the view of the Zimm plot which sometimes makes it easier to visually evaluate how well the data has been fit.

## Slide 8: Historical Data Analysis
Step 1:
Light Scattering Equation (Zimm formalism)
Step 2:
Take low angle limit,   0°, P(0°)  1
Step 3:
Take low concentration limit, c  0
Known quantities: R(), K*, c
Want to know: M, rg, A2
Notes: Notes:
The Rayleigh-Gans-Debye (RGD) equations can be recast to facilitate data analysis using the Zimm formalism.
In their new form, the light scattering equations offer a simple means to retrieve the quantities of interest from linear relationships at low angle and concentration limits. For example, in the low angle limit the form factor P(0) = 1. It is then easy to write the Zimm formalism equations such that the second virial coefficient can be obtained from the slope of a linear relationship, and the molar mass from the intercept.
Similarly, the low concentration limit equation can be written in a linear form.. If the expression for the form factor P() is written out (and a little algebra performed!), the slope of the resulting linear equations gives the rms radius, and the intercept the molar mass.
Key Ideas
Zimm formalism - In the Zimm formalism, the RGD equations are recast such that K*c/R() is on the left-hand side of the equations. Recasting the equations this way makes it possible to obtain the quantities of interest from the data using simple linear extrapolations in a Zimm or Debye plot (to be discussed later). As will be discussed in tomorrow’s lecture, there are also Berry and Debye formalisms that are useful for different polymer size ranges.

## Slide 9: The Zimm Plot Combines these two Concepts:
A2 from initial slope
rms radius from initial slope
Global Fit of the data:
Zero angle:
Zero concentration:
Notes: Notes:
What makes the Zimm plot unique is that both the angular and the concentration dependence of the scattering is displayed in one plot. From a single Zimm plot the molar mass M, the rms radius rg, and the second virial coefficient A2 can be determined.
The rms radius is always obtained from the angular dependence of the light scattering. It is computed from the slope of the angular dependence after a fit to zero concentration.
The second virial coefficient is always obtained from the concentration dependence of the scattering. It is computed from the slope of the concentration dependence after a fit to zero angle.
Key Ideas
Zimm plot - A graphical representation of the data that enables the molar mass, rms radius, and second virial coefficient to be retrieved from light scattering data. To make a Zimm plot, it is necessary to make light scattering measurements at multiple angles and concentrations. In tomorrow’s lecture the Debye plot will be discussed, which is a graphical method to retrieve the molar mass and rms radius by measuring the light scattering at multiple angles with only one concentration.
Stretch factor k - A factor determined by the ASTRA software to put the values of kc and sin2(q/2) in the same numerical range for a clearer visual presentation of the data. The final results are independent of the stretch factor.