## Slide 1
Specific Refractive Index Increment
&
Concentration Detectors
Light Scattering University
Wyatt Technology Corporation
Santa Barbara, California
Notes: Notes:
Additional information for each slide is presented in this section.
Please feel free to use this space to take notes.

## Slide 2: Topics
Specific Refractive Index Increment, dn/dc
Methods to Obtain dn/dc:
  Measuring dn/dc off-line (batch)
  Measuring dn/dc on-line (SEC)
Determination of the Calibration Constant for analog voltage Refractive Index detectors.
Lab practical: measure the dn/dc of a solute in solution
Notes:   Notes:
  RI detectors are universal concentration detectors, i.e. they are not structure specific.
  UV detectors are specific concentration detectors; i.e. the detector response only occurs if there are molecular structures that would absorb light at the wavelength being used.
  The combination of both UV and RI is a powerful technique to quantify subcomponents in protein conjugates or co-polymers. More information about protein conjugate analysis will be provided in the protein conjugate workshop at LSU.
  The Polymer Handbook and Huglin’s text have extensive lists of dn/dc values. However, they may not list a dn/dc value for your polymer, in your solvent, at your wavelength. See also www.ampolymer.com for a list of dn/dc values.
  Your USB memory drive also contains values for dn/dc: LSU Training Binder\12 - Measuring dndc

## Slide 3: Specific Refractive Index Increment dn/dc
dn/dc describes the change of the refractive index of a polymer solution with change of the polymer concentration.
dn/dc has units of ml/g; it expresses how much the aRI of a solution would increase for every g of solute added to 1 ml of solvent.
The dn/dc of a polymer solution must be measured as the polymer concentration approaches zero.
dn/dc needs to be determined:
  for your polymer,
  in your solvent,
  at the wavelength of your LS detector,
  to accurately determine the molar mass by light scattering, even when using UV as a concentration detector!
Notes:   Notes:
  A solvent by itself does not have a dn/dc; it has an absolute RI value. It is only when some solute is added to a solvent that you have a dn/dc value that expresses how the RI of the solution changes as more solute is added to solvent. For example: the dn/dc of the polysaccharide dextran in water is 0.138 ml/g at 658 nm. The absolute RI of water is approximately 1.330. If you added 1 g of dextran to 1 ml of water, the refractive index of the solution can be calculated to be: 1.330 + 0.138 = 1.468.
  The dn/dc value is one of the parameters in K* required to compute molar mass from light scattering data.
  The dn/dc value is one of the values used to compute the concentration of the polymer from the net RI signal.
  The effect of temperature on dn/dc is negligible.

## Slide 4: Comparison of dn/dc values for Different Polymer-Solvent Systems
Chemical | Refractive Index (n) | dn/dc of PS [mL/g] at 633 nm
Polystyrene (PS) | 1.59 | XXX
Toluene | 1.497 | 0.11
THF | 1.407 | 0.185
Trichlorobenzene | 1.572 | 
0.02 (measured: 0.04)
The refractive index increment can be estimated as*:
dn/dc  p(npolymer – nsolvent)
or:
(dn/dc)solvent B  (dn/dc)solvent A – p(nsolvent B – nsolvent A)
p = partial specific volume, usually set to unity
*from: “Size Exclusion Chromatography”, by S. Mori, H.G. Barth, Springer Verlag, 1999
?
Notes: Notes:
The dn/dc value for a block copolymer is the weighted average of the dn/dc values for both homo polymers:
  (dn/dc)copolymer = WA(dn/dc)A + WB(dn/dc)B
WA, WB: weight fractions of the A and B units in the copolymer.
The dn/dc of a random copolymer is not as straightforward to calculate, and should probably be measured directly.

## Slide 5: Wavelength Dependence of dn/dc
dn/dc typically increases as the wavelength of measurement decreases. Generally, in an aqueous solvent, the difference between the dn/dc values at 658 nm and 940 nm is 5% or less.
If you have dn/dc values at several wavelengths, curve fitting with extrapolation or interpolation can be used to estimate the dn/dc value at the desired wavelength.
For Mw calculations in ASTRA, it is best to use the dn/dc value for your polymer at the laser wavelength of your LS detector!
488 nm
633 nm
690 nm
Notes: Notes:
The values at 690 nm, 633 nm and 488 nm were measured with a Optilab DSP.
  The gallium arsenide laser used in HELEOS or TREOS operates at 658 nm.
  The gallium arsenide laser used in an EOS or miniDAWN operates at 690 nm.
  The wavelength of the Helium Neon gas laser in the DAWN DSP series of instruments operates at 633 nm.
  The argon ion laser option that was available for the DAWN DSP operates at 488 nm.
  The Neodymium-YAG laser (option) operates at 532nm in the EOS.
Additional dn/dc values for NaCl in H2O are from : A. Becker, et al., Berichte der Bunsengesellschaft fur Physikalische Chemie, April 1995, Vol. 99(4), pp. 600-608.

## Slide 6: Molar Mass Dependence of dn/dc
  EXAMPLE: dn/dc values of PEG in Water
dn/dc of PEG in water approaches a constant value at 10,000 Da molecular weight.
Notes:     Notes:
    Above a certain molar mass, the dn/dc can be considered to be a constant.
      dn/dc = d + e/Mw
      d, e are constants; d corresponds to the specific refractive index increment at infinite Mw
    References:
    “Size Exclusion Chromatography”, by S. Mori, H.G. Barth, Springer Verlag, 1999
    “Light Scattering from Polymer Solutions”, by Huglin, page 246
    Good Reference for PS and isobutylene: “Characterization of Low-Molecular-Weight Polymers: Failure of Universal Calibration in SEC”, R.R. Chance et. al., Int. J. Polymer Analysis & Characterization, 1995, Vol. 1, pp. 3-34.

## Slide 7: Measuring dn/dc (off-line)
Prepare 6 or more defined concentrations of your polymer in your mobile phase.
A large enough sample volume must be injected to get a flat region of constant concentration for each concentration standard.
Always inject pure solvent first and last which will be used for setting baseline.
solvent
Flat plateaus of known concentrations
Notes: Notes:
See also Chapter 6 of the Optilab rEX Hardware Manual for additional hints on dn/dc measurements.
Polyelectrolytes (polymers with charges along their backbone) should be dialyzed against the desired mobile phase prior to off-line dn/dc determination.
ASTRA software collects the voltages, performs the calculations, plots the data, and reports the dn/dc value. If you have an Optilab rEX, the data can be collected digitally into ASTRA. With other RI detectors, including the Wyatt Optilab DSP, the data will need to be collected as an analog voltage on one of the aux channels of your light scattering detector, and you will need to provide the RI calibration constant, which is the dn/dV (discussed later in this section).
When using Wyatt’s dndc kit and an HPLC pump and mobile phase degasser, a small difference in DRI will be visible for degassed mobile phase and non-degassed solvent injection. Draw the baseline through the plateaus of the injected solvent which has the same amount of dissolved gas present.
It is possible to use both a DAWN or MiniDAWN and an Optilab to simultaneously perform a Zimm Plot and dn/dc measurement (using polymer solutions of known concentration). See:
Imad A. Haidar Ahmad and Andre M. Striegel, A Coupled MALS-DRI Method for Simultaneous Zimm and dn/dc Plot Construction Instrumentation Science and Technology, 37: 574–583, 2009

## Slide 8: Measuring dn/dc off-line
The Optilab rEX measures Δn for each concentration, and the ASTRA software calculates the best fit to the Δn vs. concentration data.
Plot Δn vs. csample: The slope equals d(Δn)/dc or dn/dc
Notes: Notes:
The only user input needed to determine dn/dc is to draw the baseline, set peak limits, and provide the concentration of each selected peak. ASTRA will plot the data and determine the slope, which is the dn/dc of your polymer in your solvent.
Error propagation of dn/dc when using RI as a concentration detector:
  Recall:
  With an RI,
→ Therefore, in SEC-MALS, using an RI detector as the concentration detector, an error of 5% in the dn/dc value will result in an error of 5% in the computed molar mass.

## Slide 9: Measuring dn/dc On-line Using SEC/RI
ASTRA will compute a dn/dc value from an SEC or FFF chromatogram if you know:
  Starting concentration of your polymer sample,
  Sample volume injected,
  exact flow rate through the detector,
  100% of the injected polymer elutes from column or channel.
  Note: If you are using a analog RI detector, you must know its calibration constant (i.e., dn/dV).
Notes: Notes:
An on-line dn/dc value is equivalent to an off-line dn/dc value from a dialyzed sample.
It is generally fair to assume 100% mass recovery of organic polymers for SEC in organic solvents. It is not a safe assumption when analyzing proteins or other polyelectrolytes in aqueous buffers, however.
If you work with polymers that absorb UV light, and you know the polymer’s extinction coefficient, you can use an on-line UV detector to calculate the amount of polymer recovered in a particular peak. You can then enter that measured value as the amount of mass of sample injected, and then process the data using the RI signal assuming 100% mass recovery instead of providing either a known dn/dc value or a known RI aux constant.

## Slide 10: Analog RI Detector Calibration
The calibration constant RICC is the proportionality constant between the difference in refractive index (Δn) and the change in the voltage (ΔV) of an analog RI detector:
ASTRA uses the analog voltage (i) produced by the RI detector, the RI calibration constant (RICC), and the dn/dc value to compute the polymer concentration in a data slice.
Notes: Notes:
The Optilab rEX can send digital data directly to ASTRA 5, and it automatically calculates the differential refractive index in a data slice. All other RI detectors, including the Optilab DSP, output analog voltages (V). ASTRA collects these voltages through one of the DAWN or miniDAWN light scattering detector’s auxiliary channels and uses the RI detector calibration constant to calculate the dRI value.
The dRI calibration constant for the Optilab TrEX or rEX when collecting data in “digital” mode (using Ethernet connection) is different than an “analog” RI detector calibration constant. For an Optilab TrEX or rEX, the dRI calibration constant in “digital” mode relates light beam movement to the dRI signal. Its value is in RIU/pixel. This approach allows the instrument’s dynamic range to exceed 21 bits of resolution when transmitting data to ASTRA digitally.
Collecting “analog” voltage data from the Optilab TrEX or rEX will reduce its overall range, and the calibration constant will reflect the RIU per volt ratio (i.e., dn/dV). Please see the Optilab rEX hardware manual for additional details.
The calibration constant (RIcc) of either the Optilab TrEX or rEX, Optilab DSP, or any analog RI detector is an instrument constant, and is therefore independent of the solvent used for off-line or on-line measurements. The RIcc value measured using a series of NaCl in water solutions is applicable to any mobile phase used in later measurements, regardless of solvent. No recalibration is required.

## Slide 11: Analog RI Detector Calibration Procedure
Prepare calibration standard solutions of known concentrations (c) and known dn/dc (e.g., NaCl in water).
Pass each standard through the RI detector and record the RI detector voltage (V).
A large enough sample volume must be injected to get a flat region of constant concentration for each standard.
Concentrations from 0.1 to 1 mg/mL NaCl are appropriate for a standard analog RI detector.
Notes: Notes:
As in dn/dc determinations, always inject pure solvent before and after all samples to draw a baseline.
Determining the dRI and aRI calibration constants of an Optilab rEX :
  ASTRA 6: File  New  Experiment from method  System methods  RI Measurement  RI calibration. After calibration, you can upload the new calibration constant to your instrument by clicking on the Upload to Instrument button in the RI Calibration procedure.
  ASTRA 5: File  New  Experiment from Template  System Templates  RI Measurement  RI calibration
  Measuring an analog RI detector calibration constant will require the use of a light scattering detector that can capture the analog voltage output from an analog RI detector through an aux channel and convert it to digital information with its built-in analog-to-digital converter.
    ASTRA 6: Apply the RI calibration method in System methods  RI Measurement to your default method.
    ASTRA 5: Replace the configuration in the template above with: Example Configurations  RI Measurement  generic RI {your MALS detector}  LS batch

## Slide 12: Analog RI Detector Calibration using ASTRA
ASTRA will compute the n for each standard solution from:
n = c x (dn/dc)
Plot n vs. V: the slope is the analog RI calibration constant (AUX calibration constant needed for ASTRA):
Notes: Notes:
Standard solutions were passed through an analog RI detector, using an HPLC pump at 0.6 mL/min (no columns in the system) and a manual injection valve with a 1.0 mL sample loop.
Samples can also be passed through the RI detector using syringes and a syringe pump. (Avoid air bubbles) An HPLC pump and autosampler can be used if the volume of the sample loop in the autosampler is sufficiently large to achieve a plateau voltage. In general, a minimum of 0.5 mL is needed for this purpose.
An example data file for a RI calibration with a generic RI can be found on your USB memory drive in LSU Training Binder\12 - Measuring dndc.

## Slide 13: Conclusions
Knowledge of dn/dc is necessary for:
  Molar mass determination from light scattering measurements,
  Determining concentrations by an RI detector,
  Analog RI calibration.
The value of dn/dc depends on:
  polymer composition,
  solvent,
  molar mass,
  laser wavelength.
The dn/dc can be obtained from:
  the literature,
  measured directly with an RI detector (online or offline)
For greatest accuracy, the dn/dc used in light scattering experiments should have been obtained for your polymer, in your solvent, and at the wavelength of your LS detector.

## Slide 14: Lab Practical:
This exercise will measure the dn/dc of an solute in solution (NaCl in H2O) using ASTRA.
Required Materials:
  Calibrated RI detector
  Several solutions of known concentration
  Computer with ASTRA installed
Helpful Accessories:
  Syringe pump and syringes or:
  Wyatt DNDC injector kit
  HPLC pump and injection system with sample loop of 0.5 ml or greater
Notes: Notes:
ASTRA 6:
  Configuration Wizard  Batch  Select your Refractometer (Optilab rEX/TrEX or generic)  Name Configuration  OK. This configuration will determine dn/dc from your RI instrument. Then select File  New  Experiment from Default, select 30 min for the Collection Interval and start the experiment.
  Alternatively: File  New  Experiment from method  System methods  RI Measurement  batch (determine dndc)., set Collection Interval and start the experiment
ASTRA 5: Create a new Experiment from Template:
  File  New  Experiment from Template System Templates  RI Measurement  batch (determine dndc)
  Experiment  Configuration
     Optilab rEX  Select Physical Instrument
     Solvent: water
  Experiment  Procedures  Set Collection Interval to 30 min (use “Run Indefinitely”, if needed). Start the experiment.
Inject solvent for baseline, 5-6 different sample concentrations, solvent for baseline again;
When finished:
  Draw baselines
  Select peaks of plateaus (flat region) and enter concentration for each injection
  Click on dn/dc from RI to see the fit and results for dn/dc. Disable peaks if needed.
An example ASTRA file is located on your USB memory drive in: LSU Training Binder\12 - Measuring dndc.
Hints on how to measure dndc in organic solvents can be found on your USB memory drive in: LSU Training Binder\12 - Measuring dndc \dndc measurement strategy in organic solvents.

## Slide 15: Appendix
Use of a UV Detector as a
Concentration Detector

## Slide 16: Use of a UV Detector
A UV detector may be used as the concentration detector in HPLC SEC-MALS under these conditions:
  The extinction coefficient of the polymer at the wavelength of the UV detector must be known.
    For proteins, this value can be estimated from the amino acid composition or simply measured using a UV detector and a known concentration of sample.
    It is also possible to determine the UV extinction using ASTRA if the dn/dc value is known.
  The response factor (AU/V) of the UV detector must be known.
  The cell path length (L) of the UV detector must be known.
Notes: Notes:
The UV response factor can usually be found in your UV detector’s manual, and will be listed as absorbance units per volt (AU/V). The manual should also provide the flow cell’s path length value.
At 280 nm, the molar absorption coefficients of a peptide can be approximated by the weighted sum of the 280 nm molar absorption coefficients of the amino acids tryptophan (W), tyrosine (Y) and cysteine (C) by the following relationship:
  ε = (nW×5500) + (nY×1490) + (nC×125)
  with n being the number of each amino acid residue.
See also:
  Gill, S.C. and von Hippel, P.H. (1989). Calculation of protein extinction coefficients from amino acid sequence data. Anal. Biochem. 182:319-26.
  Pace, C.N., et al. (1995). How to measure and predict the molar absorption coefficient of a protein. Protein Sci. 4:2411-23.
Values for UV extinction coefficients from the literature are also compiled in:
Practical Handbook of Biochemistry and Molecular Biology, Fasman, D.G., Ed. (1992). CRC Press, Boston.
See also Pierce (www.piercenet.com) Tech Tip #6 for a useful discussion on extinction coefficients.

## Slide 17: Concentration from a UV Detector
The quantity recorded by the UV detector is the absorbance A:
   is the extinction coefficient in mL/ (g x cm),
  c is the concentration in g/mL,
  L is the detector path length in cm.
ASTRA converts the UV voltage into absorbance A using the known UV response factor and cell path length.
The absorbance is converted into concentration based on the known extinction coefficient.

## Slide 18: SEC-MALS of BSA using UV detection
Example: information that would be needed to enter in ASTRA for determining Mw and recovered mass of Bovine Serum Albumin (BSA) with a MALS detector and UV detector at 280 nm:
    dn/dc = 0.185 mL/g (at 690 nm)
    g/mL = 660 AU/(cm x g/mL) at 280 nm
    UV cell path length = 0.60 cm
    UV response factor = 1 AU/V
Note that even if a UV detector is used for on-line sample concentration detection, it is still necessary to know dn/dc for the light scattering equation.

## Slide 19: Using an RI and a UV Detector
If both an RI and a UV detector are used with MALS, the dn/dc value of a polymer can be measured on-line.
  The actual mass (g) of the polymer eluted can be computed in ASTRA using the data from the UV detector as long as the  is known.
  Using this calculated mass as the ‘injected mass’, the RI signal can then be used to compute the dn/dc value.
  The use of the UV detector to calculate the real recovered mass eliminates the assumption that 100% of the polymer elutes from the SEC column.
When the dn/dc value is known, the combination of RI and UV detectors can also yield the extinction coefficient of the peak.
When dn/dc and the extinction coefficient are known, the combination of RI-UV-MALS can yield information about copolymers and modified proteins: calculated % pegylation, weight % of each component, etc.
Notes:   Notes:
  The dn/dc value is required to compute K* in the light scattering equation.
Note that if both a UV and an RI detector are used, the extinction coefficient can be measured on-line.
The actual mass of the polymer eluted can be computed in ASTRA using the RI detector, when dn/dc is known. Using this mass as the ‘injected mass’, the UV signal can then be used to compute . This is a relatively crude way to measure , but it can be useful.
More about the analysis of modified proteins and copolymers in the Protein Conjugate Workshop on the third day of LSU.