## Slide 1
Optilab rEX and TrEX
Light Scattering University
Wyatt Technology Corporation
Santa Barbara, California
Notes: Wyatt Technology’s Refractometer Extended Range
4 Topics in this talk
General Method of Operation
How dRI is measured
dRI and aRI modes
Daily use

## Slide 2: What is RI?
Refractive Index: the ratio (n) of the velocity of light in a vacuum (c) to that in a medium (v)
n = c/v

## Slide 3: Snell's Law: n1 sin(Θ1) = n2 sin(Θ2)
Scenario 1:
Light strikes the new medium “normal” to the plane
sin(Θ1) = 0
NO refraction
Medium 1
Medium 2

## Slide 4: Snell's Law: n1 sin(Θ1) = n2 sin(Θ2)
Scenario 2:
Light strikes the new medium at some non 0 angle relative to “normal”
sin(Θ1) ≠ 0
and therefore refraction
Medium 1
Medium 2
Θ1
Θ2

## Slide 5: Snell's Law: n1 sin(Θ1) = n2 sin(Θ2)
Scenario 3:
Light strikes the interface at some non 0 angle relative to “normal” but n1 = n2
No Refraction
Medium 1
Medium 2

## Slide 6: Scenario 1 and 2 Combined
Medium 1
Medium 2
Medium 1

## Slide 7: The Sample Cell
Medium 1
Medium 2
(Sample)
Medium 1

## Slide 8: The Sample and Reference Cell
Medium 1 (Sample)
Medium 2 (Reference)
When Sample is the same as the Reference: No refraction.
When Sample is different from the Reference: Refraction!

## Slide 9: Traditional Split Photodiode
Traditional method (the competition):
Notes: Previous refractometers built around the generated angle.
When s=r no refraction and the exiting beam
fell upon 2 detector photodiodes.
Both photodiodes generated voltages
When s was not= to r, the beam exited at an angle
Causing the beam to ‘walk off’ it’ normal place on the detectors
changing the voltages generated by the photodiodes.
By calibrating the instrument, the voltages changes would be
correlated to Angle of Refraction, but user had to attenuate the
light source to keep the beam from missing the detectors
User had to choose between sensitivity and analytical range
That’s how everyone else does it.

## Slide 10: Photodiode Array
The Wyatt Way:
Differential RI Calibration – RIU/pixel
(internal constant not shown in ASTRA)

## Slide 11: Absolute Refractive Index

## Slide 12: Absolute Refractive Index
Optilab rEX must be in Purge Mode for aRI measurements
Solution’s aRI is shown on the rEX front panel – no need for ASTRA
aRI can be re-calibrated using three or more standard liquids, e.g. Water, Methanol, Toluene, THF

## Slide 13: LED Settings
Light Intensity must be less than 7.8V for proper operation!
Light Intensity will change significantly when solvent is changed, and will drift as LED ages (years?)
Use “Optimize” button to get close, but fine tune by hand

## Slide 14: Installation “Pointers”
30-40 minute instrument warm-up / temperature equilibration time.
Nitrogen needed only for sub-ambient operation. 50psi bleed.
DHCP Ethernet communications recommended, static IP is OK.

## Slide 15: Daily Use “Pointers”
Never leave salt solution stagnant in cell to prevent precipitation;
Never transition directly from salt solution to organic solvents;
Place the rEX last in the flow path to keep backpressure low;
Purge cell daily for ~2 minutes;
Check LED intensity periodically, especially after changing solvents;
Store in ethanol to prevent organisms from growing;
Contact Wyatt Support for help replacing the LED.