## Slide 1
Characterizing Protein Conjugates Using MALS-UV-DRI Detection Following SEC Separation
Light Scattering University
Wyatt Technology Corporation
Santa Barbara, California

## Slide 2: Protein Conjugate Analysis in ASTRA V
Analyze glycosylated and PEGylated proteins, and membrane protein-detergent complexes to determine protein fraction.
Unprecedented view of protein modification processes when coupled with band broadening correction.

## Slide 3: Theory of MALS-UV-RI Detection
For an unmodified protein, ε can be determined from an SEC run with known dn/dc, and vice versa.
From DRI:
DRICC: DRI calibration constant
I DRI: Intensity of the DRI signal
From UV:
UVRF: UV response factor
I UV: Intensity of the UV signal

## Slide 4: Theory of MALS-UV-RI Detection
For modified proteins (such as membrane protein-detergent complex, pegylated and glycosolated proteins):
x protein is the weight fraction of protein
x protein as well as molar masses of core protein and protein-modifier complex can be determined.

## Slide 5: Glycosylated Proteins
Protein X Expressed from Mammalian or Insect Cell
UV trace, Mammalian Cell
UV trace, Insect Cell

## Slide 6: Same Protein but Different Degrees of Glycosylation
Protein MW
Carbohydrate MW, Mammalian Cell
UV trace, Mammalian Cell
Carbohydrate MW, Insect Cell
UV trace, Insect Cell

## Slide 7: PEGylated Proteins
Modeling of Protein X-20k PEG Conjugate
Courtesy of
Dr. A. Ghose, Amgen
PEG (polyethylene glycol):
nontoxic, nonimmunogenic, highly water soluble, readily cleared from the body, approved by FDA.
PEG-proteins:
enhanced solubility, sustained absorption, reduced immunogenicity and proteolysis, reduced renal clearance, optimized biodistribution, reduced toxicity, improved efficacy.

## Slide 8: Degree of PEGylation is 0.5 for Protein Y
Protein Y Mw=18.2 kD
Branched PEG
Mw=18 kD
Conjugate Mw = 34 kD Degree of PEGylation = 0.5

## Slide 9: Rh Can Be Measured by On-Line DLS
Protein Y (18.2 kD)
Rh = 2.0 nm
QELS & AUX detector
LS Slice : 2042
QELS Slice : 1018
Volume : 18.613 mL
Temperature : 28.9 °C
Rh : 1.9 ± 0.1 nm
0.60
0.80
1.00
1.20
1.40
1.60
1.0x10
-7
1.0x10
-6
1.0x10
-5
1.0x10
-4
1.0x10
-3
1.0x10
-2
1.0x10
-1
1.0x10
0
1.0x10
1
Correlation function
Delay time (sec.)
QELS Plot - 05304801
Branched 18k PEG Rh = 3.7 nm
QELS & AUX detector
LS Slice : 1550
QELS Slice : 154
Volume : 15.333 mL
Temperature : 27.8 °C
Rh : 3.7 ± 0.1 nm
0.60
0.80
1.00
1.20
1.40
1.60
1.0x10
-7
1.0x10
-6
1.0x10
-5
1.0x10
-4
1.0x10
-3
1.0x10
-2
1.0x10
-1
1.0x10
0
1.0x10
1
Correlation function
Delay time (sec.)
QELS Plot - 05304601
Conjugate
Rh = 4.7 nm
QELS & AUX detector
LS Slice : 1492
QELS Slice : 744
Volume : 14.947 mL
Temperature : 29.0 °C
Rh : 4.6 ± 0.1 nm
0.60
0.80
1.00
1.20
1.40
1.60
1.0x10
-7
1.0x10
-6
1.0x10
-5
1.0x10
-4
1.0x10
-3
1.0x10
-2
1.0x10
-1
1.0x10
0
1.0x10
1
Correlation function
Delay time (sec.)
QELS Plot - 05301301

## Slide 10: Membrane Proteins
Membrane protein is soluble only in the presence of micelles.
Membrane protein – detergent complex has very different elution/migration properties than the globular standard proteins.
Molar mass of detergents in the complex needs to be measured as well in order to determine the oligomeric form of the core membrane protein.
Mass Spec often breaks down the oligomeric form.
Artist’s rendition of membrane protein in detergent/micelle solution
free micelle
membrane protein
detergent
Problems with Traditional Techniques (such as SEC with column calibration, gel electrophoresis, mass spec) in determining Oligomeric Form of Membrane Proteins.

## Slide 11: SEC-MALS Conditions
HPLC system: Agilent 1100
Columns : Shodex Protein 803 KW
UV detector: Agilent 1100 UV @ 280 nm
MALS detector: DAWN EOS from Wyatt
RI detector: Optilab from Wyatt
Sample concentration: 1 mg/mL
Injection volume: 50 µL
Membrane Protein (MP)
Expected molar mass for monomer: 48 kD
Extinction coefficient (1%): 16.6 (measured), 20 (calculated in 6M GuHCl)
Mobile Phase:
10 mM HEPES + 50 mM NaCl + 10% glycerol + 0.08% LDAO (> CMC)
(n-dodecyl-N,N-dimethylamine-N-oxide), pH 7.0

## Slide 12: Membrane Protein Z is a Monomer
BSA monomer, 67 kD
Complex, 84 kD
Detergent, 35 kD
Core membrane protein, 49 kD

## Slide 13: Calculated MM of MP as a Function of Core Protein’s Extinction Coefficient and Detergent’s dn/dc Value
dn/dc of detergent [mL/g] | e 1% of core protein | Protein % (by wt) | MM protein [kD] | MM complex [kD]
0.135 | 16.6 | 73 | 57 | 78
0.135 | 18.3 | 65 | 52 | 80
0.135 | 19.5 | 60 | 48 | 81
0.135 | 20.0 | 58 | 47 | 81
0.12 | 19.5 | 57 | 48 | 85
0.15 | 19.5 | 62 | 48 | 78
The dn/dc value for core protein is 0.185 mL/g; extinction coefficient for detergent at 280 nm is zero.